DESIGN FOR A SIMPLIFIED COCHLEAR IMPLANT SYSTEM

Soon Kwan An1,2, Se-Ik Park1,2, Choong Jae Lee1, Kyung Min Byun1, Kyu Sik Min1,2, Jung Min Lee2, Blake S. Wilson3, Dewey T. Lawson3, Stephen J. Rebscher4, Seung Ha Oh1 and Sung June Kim1

1Nano Bioelectronics and Systems Research Center, Seoul National University, Seoul, Korea

2Nurobisys Corporation, Seoul, Korea

3Center for Auditory Prosthesis Research, RTI International, NC, USA

4Department of Otolaryngology, University of California, San Francisco, CA, USA

A simplified cochlear implant system would be appropriate for widespread use in developing countries. Here, we describe a cochlear implant that we have designed to realize such a concept. The system implements 8 channels of processing and stimulation using the Continuous Interleaved Sampling (CIS) strategy. A generic Digital Signal Processing (DSP) chip is used for the processing, and the filtering functions are performed with a Fast Fourier Transform (FFT) of a microphone or other input. Data derived from the processing are transmitted through an inductive link using Pulse Width Modulation (PWM) encoding and Amplitude Shift Keying (ASK) modulation. The same link is used in the reverse direction for backward telemetry of electrode and system information. A custom receiver-stimulator chip has been developed that demodulates incoming data using pulse counting and produces charge balanced biphasic pulses at 1000 pulses/s/electrode. This chip is encased in a titanium package that is hermetically sealed using a simple but effective method. A low cost metal-silicon hybrid mold has been developed for fabricating an intra-cochlear electrode array with 16 ball-shaped stimulating contacts.

This work was supported in part by the Korea Science and Engineering Foundation (KOSEF) through the Nano Bioelectronics and Systems Research Center (NBS-ERC) of Seoul National University (#R11-2000-075-01001-0) and in part by the Nurobiosys Corporation.

Electrode development was conducted in collaboration with the Epstein Laboratory at the University of California, San Francisco with support from NIH contracts #NO1-DC-2-1006 and NO1-DC-3-1006.