A comparison of platelet function in cold stored whole blood and platelet concentrates.

Authors: Sian Huish1, Laura Green2,3,4, Carly Kempster1, Peter Smethurst1, Michael Wiltshire1, Chetankumar Prajapati5, Elisa Allen5, Rebecca Cardigan1,6

Affiliations:

1NHS Blood and Transplant, Cambridge

2NHS Blood and Transplant, London

3Blizzard Institute, Queen Mary University of London, London

4Barts Health NHS Trust UK, London

5NHS Blood and Transplant, Stoke Gifford

6Department of Haematology, University of Cambridge, Cambridge, UK.

Conflicts of interest: Rebecca Cardigan is a member of the Terumo BCT scientific advisory board.

No Ethics required.

Grants: This work was part-funded by London Air Ambulance and Barts Charity.

Contacting author: Sian Huish, NHS Blood and Transplant, Long Road, Cambridge, CB2 0PT, UK sian.huish@nhsbt.nhs.uk

Word count: 3498

Number of figures: 5

Number of tables: 2

Number of references: 36

Running title: Platelet function in whole blood

Abstract

Background:

There is renewed interest in the use of whole blood (WB) for the resuscitation of trauma patients. Platelet function in stored WB compared to platelet concentrates is not well established and was assessed in vitro in this study.

Methods:

Leucocyte-depleted cold stored WB (CS-WB) was prepared using a Terumo WB-SP Imuflex kit and held at 2–6 ºC alongside: A) UK standard pooled platelets stored at 20-24 °C (RT-PLTS) B) pooled platelets stored at 2-6 °C (CS-PLTS), C) platelet-rich plasma produced using the Terumo kit (CS-PRP), for 21 days. A series of in vitro assays assessed platelet function.

Results: Platelet count was retained to 57±14 % of starting number at day 21 in CS-WB. Over time CS-WB platelets become more activated, with increased CD62P expression (Day 1: 7±3.7 % vs. Day 21: 59±17.1 %) and annexin V binding (Day 1: 2±0.2 % vs. Day 21: 21±15.1 %). For comparison, 18.6±6 % of platelets in RT-PLTS demonstrated CD62P expression at day 7 whilst Annexin V binding in RT-PLTS at day 7 was 2.6±0.5 %. Over storage, aggregatory response to agonists decreased in all arms. Functional platelet microparticles increased steadily in CS-WB throughout storage.

Conclusion: During storage, platelet count reduced in CS-WB, whilst CD62P expression and annexin V binding increased. This was accompanied by reduced aggregatory response, although compared to 7-day old RT-PLTS, CS-WB maintained maximal response to agonists for longer, suggesting that the shelf life for CS-WB can be considered for up to 21 days.

KEYWORDS: Blood Component Preparations, Haemostasis, Whole Blood, Major Haemorrhage, Cold platelets

Introduction

There is a resurgence of interest in the transfusion of cold-stored whole blood (CS-WB) for the management of major haemorrhage in trauma, as this is likely to have a longer shelf-life than fresh whole blood transfused within 24 hours of donation. The duration of CS-WB storage currently permitted by regulators in the USA was based mainly on the viability of red cells, i.e., day 21 in citrate-phosphate-dextrose anticoagulant (CPD).1,2 However, the current interest in the use of CS-WB to provide a source of plasma, platelets and red cells in a 1:1:1 ratio, which may provide clinical , has led to renewed consideration of the haemostatic potential of this component throughout storage.3,4,5,6,7

In the past two decades, the introduction of universal leucocyte depletion (LD) by many blood operators, has meant that most whole blood collections are platelet-depleted, as the LD filters also remove platelets. An alternative platelet-sparing LD filter is available in the IMUFLEX-WB-SP collection set from Terumo, which permits approx. 80 % of platelets to pass through the filter.8

Prior studies have assessed plasma LD and non-LD cold-stored WB (CS-WB). They have shown that there is significant loss of certain coagulation factors during storage, although the ability to generate a clot in vitro is maintained well until day 21 in global assays, such as rotational thromboelastometry/thromboelastography (ROTEM/TEG) and thrombin generation. 9, 10, 11, 12, 13

Recently, studies have specifically addressed the platelet functionality in CS-WB and, separately, cold platelet concentrates (CS-PLTS).  Studies evaluating CS-WB in vitro have suggested an increase in platelet activation, a decline in platelet count, and a loss of aggregatory response over storage.12, 13 Further platelet studies suggest that CS-PLTS maintain their response to agonists for longer, have decreased metabolism, increased activation and promote superior clot formation compared to standard room temperature platelets (RT-PLTS). 14, 15, 16 Assays of platelet function are notoriously variable, making it difficult to compare data across studies/groups. Detailed assessment of platelet function in CS-WB compared with room temperature platelet concentrates, the current standard of care, are lacking. Therefore, we undertook analysis of platelet function in CS-WB compared with RT-PLTS, CS-PLTS, and a similar dose of platelet to that in CS-WB stored in the absence of red cells (i.e. platelet-rich plasma [CS-PRP]). This work was undertaken as part of a series of studies to inform an appropriate shelf life of LD CS-WB in the UK, which may be used in the future.

Materials and methods

The study design is shown in Fig. 1. CS-WB was compared to standard NHSBT pooled buffy coat platelets in SSP+ held in current storage conditions (22±2 °C, agitation, RT-PLTS) or refrigeration (4±2 ° C, static, CS-PLTS). An additional comparator, platelets in plasma from a single LD whole blood donation (cold stored platelet rich plasma [CS-PRP]), were also used to understand the effect of storing CS-WB platelets at 4 °C in the absence of red cells. A total of 10 units of CS-WB were also stored at 2–6 °C for 14 days without manipulation (CS-WB (NM), not included in Fig. 1), before being tested to assess how periodic warming and mixing required to sample units affected measured results in CS-WB. Owing to the different platelet concentration and storage media, not all arms could be compared with each other in all assays. This is indicated in the methods below and the results.

Blood processing and sampling

A total of 110 units of whole blood from donors, 475 ± 47.5 mL, were collected into 66.5 mL CPD anticoagulant, 80 in bottom-and-top (BAT) collection packs (LQT614B, Macopharma, Twickenham, UK) and 30 in top-and-top (TAT) collection packs (FQE6283LB, Macopharma) and stored at 22±2 oC for up to 24 hours. TAT units were filtered using the Imuflex WB-SP platelet-sparing filter kit (Terumo BCT, Larne, UK), attached via sterile docking using a Terumo Sterile Connecting Device (TSCD, Terumo BCT). A total of 10 units were CS-WB and 10 units CS-WB (NM). The remaining 10 filtered TAT units were processed into CS-PRP by centrifugation and separation using a semi-automatic press (Luxomatic, LMB Technology, Milton Keynes, UK). BAT units were processed into standard pooled platelets in SSP+ (prepared from four donations). Two platelet units were pooled together and split, with one assigned to standard storage at 22±2 °C under agitation (RT-PLTS, the current standard of care in the UK) and the other held at 4±2 °C, without agitation (CS-PLTS).

Samples were taken from units after processing via sterile connection of a sample pouch (VSE000A, Macopharma). Units of CS-WB, CS-PLTS, and CS-PRP were then placed at 4±2 °C (without agitation) and further sampled on days 2, 5, 7, 9, 12, 14, 16 and 21 of storage. RT-PLTS were sampled after processing and placed at 22±2 °C under continuous agitation and sampled as for other arms of the study. CS-WB (NM) was sampled after processing, placed at 4±2 °C and only tested further at day 14.

Platelet counts

Platelet count (fluorescence) for all units was assessed on a sample (500 µL) incubated (RT, 30 min, roller mixer) in an EDTA tube (1.3 mL, Sarstedt, Leicester, UK) before analysis using a XN-1000 haematology analyser (Sysmex, Milton Keynes, UK).

Annexin V

CD61-positive platelet events that bound annexin V (%) were assessed by ﬂow cytometry (Gallios, Beckman Coulter, High Wycombe, UK). In brief, 5 µL of sample was incubated at RT for 20 min with 2 μL monoclonal mouse anti‐human CD61-APC antibody (BD Biosciences, Wokingham, UK), 4 µL Annexin V –PE (BD Biosciences), 5 µL GPRP (final concentration 4 mM) (Merck, Darmstadt, Germany) and HEPES‐buffered saline with calcium (final concentration 3 mM). Non‐specific binding was corrected using an APC isotype IgG1, k reagent (BD Biosciences) or EDTA (Sigma).

Platelet activation markers and response to agonists

Activation of CD42b-positive platelet events was assessed via CD62P expression determined by flow cytometry. In brief, 2 µL of sample was incubated at RT for 20 min with 2 μL monoclonal mouse anti‐human CD42b-APC antibody (BD Biosciences), 5 µL monoclonal mouse anti-human CD62P (Beckman Coulter) and HEPES‐buffered saline. Non‐specific binding was corrected using an APC isotype IgG1, k reagent (BD Biosciences). Samples were also incubated at RT for 20 min as above with a range of ADP (0.2 – 2 µM, Sigma), TRAP-6 (4 - 32 µM, Sigma) and CRP (0.1 – 2.0 µg/mL, University of Cambridge, Cambridge, UK) to determine a response.

Impedance aggregometry

Responses to ADP (6.5 µM, 30 µM Sigma) and TRAP-6 (4 µM, 32 µM Sigma) were assessed by impedance aggregometry using a Multiplate Analyser (Roche Diagnostics, Burgess Hill, UK). CS-WB was tested neat. RT-PLTS and CS-PLTS were diluted to a fixed concentration of 600 x 109/L using autologous platelet-poor plasma (PPP), prepared each testing day from the units by double centrifugation (2862 x g, 10 min). A fixed concentration of 600 x 109/L was chosen to ensure a satisfactory area under the curve (AUC), and provide a greater scope for sample dilution over storage considering the platelet count decrease over time.

Platelet microparticles and functionality

Supernatants were separated via double centrifugation (2862 x g, 10 min).

To measure microparticles by flow cytometry, 25 µL of supernatant was incubated at RT for 20 min with 5 µL monoclonal mouse anti-human CD61-APC antibody (BD Biosciences), 10 µL Annexin V-PE (BD Biosciences), and 110 μL of HEPES-buffered saline supplemented with calcium (2 mM) and GPRP (4 mM, Sigma) in a BD Trucount tube (BD Biosciences). Non‐specific binding was corrected using an isotype reagent (BD Biosciences) or EDTA (Sigma) as appropriate.

To assess the ability of supernatant from CS-WB to promote thrombin generation in phospholipid-depleted plasma a commercial kit was used (STA–PROCOAG–PPL Kit, Diagnostica Stago, Reading, UK). The dilution of plasma with SSP+ in standard room temperature platelets (RT-PLTS) and cold-stored platelet concentrates (CS-PLTS) meant it was not possible to measure activity in these units.

Lactate dehydrogenase

Supernatants were separated via double centrifugation (2862 x g, 10 min) and stored at ≤-30°C until testing for extracellular lactate dehydrogenase using an Alinity CI clinical chemistry analyser (Abbott Diagnostics, Chicago, USA).

Statistical analysis

Multivariable, linear mixed models were used to estimate and compare the change over time in outcome variables in different arms. In contrast to a linear model, a linear mixed model allows correlation between observations measured within the same experimental unit. Change over time was compared between arms and reported for Fig. 2A, 2B and 3. The models contain an interaction between time(days) and arms, and random effects at experimental unit level. The non-linear trend in outcome variables were modelled either using restricted cubic splines (RCS) (Fig. 2A-D,3,4B-D) and by segmented regression (Fig. 4A, Fig. 5A,B). The knots of RCS models, and breakpoints in segmented regression models were chosen by Akaike Information Criterion (AIC). The models were compared using likelihood ratio test. Specific comparisons between results for day 7 RT-PLTS and all components at day 14 and 21 were made to understand how platelets compared to current standard of care. Specific comparisons were made between CS-WB and CS-PLTS at days 7, 14 and 21 for data in Fig. 2C and 2D to understand differences between cold platelets held in concentrates and WB. We used linear contrast to estimate the difference between different arms at different time points and the p-values were adjusted using Tukey’s method. Analyses were conducted in R statistical software (version 4.0.2).

Results

Final composition of all arms studied are given in Table 1. Cold-stored whole blood (CS-WB) and CS-WB(NM) are grouped together (N = 20), as these were produced using Terumo Imuflex WB-SP. All other arms (CS-PRP, RT-PLTS, CS-PLTS, N= 10 for each) are described separately.

Platelet counts

Leucocyte depletion (LD) of CS-WB resulted in approximately 20% loss of platelets (data not shown). Platelet count in CS-WB decreased gradually during storage (Fig. 2A), and at days 14 and 21 the mean platelet count were 125±55 x 109/L (72±13 % of day 1 values) and 113±50 x 109/L (57 ± 14 % of day 1), respectively. When calculating the platelet count as percentage of the original donation (i.e. pre-LD filtration, fresh whole blood) for CS-WB, the platelet count was of 60±13 % at day 14, and 47±13 % at day 21. Change over time for CS-WB was significantly higher than RT-PLTS and CS-PLTS (P<0.0001), but not CS-PRP. CS-WB (NM) had a mean platelet count of 148±51 x 109/L at day 1 and 120±48 x 109/L at day 14 (80±12 % of day 1).

Lactate dehydrogenase (LDH)

Supernatant LDH increased steadily throughout storage in all components (Fig. 2B). Comparison of CS-WB with CS-PRP, platelet-rich plasma separated from CS-WB, provides an indication of the relative contributions of red cells and platelets in CS-WB where red cell lysis may also contribute towards LDH levels. The minimal difference here suggests the increase in LDH is primarily due to platelet lysis.  The dilution of plasma with SSP+ in the manufacturing of RT-PLTS and CS-PLTS means the baseline LDH values are lower than CS-WB or CS-PRP.  The change over time in supernatant LDH levels over storage in CS-WB was not statistically different compared to RT-PLTS, CS-PLTS or CS-PRP.

Annexin V binding

CS-WB and CS-PRP showed a greater increase in Annexin V binding across storage (from baseline (2 %) on day 1 to 21±7 % and 18±7 % at day 14, respectively) than RT-PLTS and CS-PLTS (which also start at baseline (2 %) and rise to 4±1% by day 14, and then to 22±16 % and 6±2 % respectively at day 21) (Fig. 2C). Annexin V binding in CS-WB is significantly greater than in CS-PLTS at days 14 and 21 (P<0.0001).

P-selectin expression

P-selectin (CD62P) was measured in resting platelets in each component over storage. As detailed in Fig. 2D, CD62P expression increased steadily in all components. By day 14, a mean value of 60±17 % was seen in CS-WB compared to CS-PRP (54±14 %) RT-PLTS (32±12 %) and CS-PLTS (50±9 %). At day 21, a mean of 59±17 % was seen for LD-WB, whilst CS-PRP, CS-PLTS and RT-PLTS were 52±8 %, 50±11.3 % and 42±11 % respectively. There was no significant difference between CD62P expression in CS-WB and CS-PLTS at days 7, 14 and 21.

Response to agonists by flow cytometry

A dose response to ADP, TRAP-6 and CRP-XL was demonstrated in all components (Fig.3 and supplementary data S1). The maximum percentage positivity of CD62P expression was largely maintained in cold-stored components to day 14. Change over time in maximal response was significantly lower in CS-WB than RT-PLTS for all three agonists (P<0.0001), and significantly different between CS-WB and CS-PLTS with CRP-XL (P=0.0008). Interestingly, cold-stored components demonstrated an increase in CD62P expression in response to CRP-XL after initial storage at 2-6 °C, indicated by increase between days 1 and 2.

Response to agonists by Impedance Aggregometry

Platelet aggregation in response to ADP and TRAP-6 decreases rapidly in CS-WB over storage (Fig. 4), correlating with the reduction in difference between maximal and baseline CD62P expression in Fig. 3. RT-PLTS and CS-PLTS aggregatory responses were also assessed at a fixed concentration. It is difficult to compare area under the curve (AUC) values between CS-WB and platelet concentrates, owing to different platelet count and suspending media. However, a loss of aggregatory response was also seen during storage of RT-PLTS and CS-PLTS (Fig. 4). CS-PLTS retained response to agonists for longer than RT-PLTS.

Platelet microparticles and functionality

Annexin V binding CD61-positive microparticles were quantified in supernatants from CS-WB. The ability to promote thrombin generation and clot formation in phospholipid-depleted plasma were assessed in a commercially available kit. Microparticle count increased steadily throughout storage in all CS-WB units (Fig. 5). As microparticles levels increased throughout storage, the time to clot decreased (Fig. 5).

Comparison of CS-WB units regularly sampled and mixed, with non-manipulated CS-WB (NM)

For LDH, annexin V binding, CD62P expression, response to ADP, TRAP-6 and CRP-XL by flow cytometry, aggregatory response to ADP and TRAP-6 and microparticle content and functionality, cold-stored whole blood (CS-WB) and non-manipulated CS-WB (NM) were comparable at day 14.

Discussion

Until the 1970’s platelets were stored at 4 °C. After that, the storage temperature was raised to 22 °C (under agitation) following studies demonstrating reduced survival with CS-PLTS vs. RT-PLTS for chronic thrombocytopenic patients. 17,18,19, 20,21. In the last 2 decades, the optimal storage temperature for platelet components administered for treatment of bleeding patients has been challenged, with studies now arguing that in bleeding patients, a reduced platelet survival of one day (vs. 5 days for RT-PLTS) and platelet function for at least 6 hours, might be sufficient to keep critically bleeding patients stable until surgery. These platelets could also be more effective in promoting haemostasis and rapidly stopping bleeding.22 This is supported by in vivo studies in normal volunteers, 17,19 preliminary results from studies in cardiothoracic patients23 and animal bleeding model24, all hinting that CS-PLTS are capable of correcting bleeding times up to 24 hours following transfusion. However, there have been just two clinical studies to assess safety of CS-PLTS vs. RT-PLTS or demonstrate efficacy in bleeding patients – a paediatric study and a pilot study in cardiothoracic surgery 25, 26.

The viability of red cells stored at 4 °C is well understood, but the consequences for platelet function when stored at 4 °C, with red cells (like CS-WB), are less well defined. Similarly, direct comparison of platelets in CS-WB and RT-PLTS are limited. With increasing interest in the use of CS-WB for resuscitation of patients, greater understanding of platelet function is required to inform a suitable shelf life for CS-WB. This work specifically focused on platelet function in CS-WB, compared with CS-PLTS, RT-PLTS and a CS-PRP, to understand differences in functionality.

Our results show that over time platelets in cold-stored components increase their basal levels of activation, whilst maintaining maximal agonist-induced response by flow cytometry. With RT-PLTS, basal levels of activation also increase, but the maximal response decreases with time. When assessing the units with impedance aggregometry, the sharp decline in response in components mirrors the decreasing gap between the basal P-selectin expression and maximal agonist response, as determined by flow cytometry. This correlates with published data assessing aggregation response to ADP and TRAP in CS-WB (Table 2), where a similar loss was observed over 15 days storage.11,12 These findings are supportive of the idea that CS-WB and CS-PLTS are hyper-aggregatory compared to RT-PLTS.11,12,13,14 In keeping with other published work where CS-WB was agitated, we did not see differences between CS-WB and CS-WB(NM).12, 27

At the beginning of storage, annexin V binding (Phosphatidylserine expression) and P-selectin expression are low in CS-WB. This indicates minimal activation of platelets, and thus a stronger response to agonists compared to platelets stored beyond 5 days. Studies have shown that P-selectin increases dramatically across 21 days of storage in CS-WB,11 and the increase is greater in CS-PLTS than RT-PLTS.15 In this study, P-selectin expression was greater in CS-WB than CS-PLTS, and maximal response to agonists was better preserved in CS-WB compared with CS-PLTS. Interestingly, Slichter et al noted that platelet in vivo recoveries are higher for a given shelf-life in CS-WB compared with CS-PLTS, suggesting that platelet function might be better preserved in CS-WB.22

Increased annexin V binding, P-selectin expression, microparticle generation and retained response to agonists, as seen in CS-PLTS and CS-WB in this study, may make these products ideal for bleeding patients, where immediate achievement of haemostasis is more important than longer-term survival of platelets. However, increased activities of the above platelet markers might pose an increase in thrombotic risk for these patients. To date there have been very few studies 25,26 that have compared the efficacy and safety of CS-PLTS or CS-WB vs. current component therapy in bleeding patients, indicating the need for well-designed large clinical studies to address this question. The relative contribution of microparticles to haemostasis in major haemorrhage is not fully elucidated. We observed an increase in platelet microparticle content in parallel with a reduction in platelet count in CS-WB during storage. There was also a concomitant increase in the ability of these MP’s to promote thrombin generation in phospholipid-depleted plasma, as noted in other studies.28, 29 However, clinical studies with cryopreserved platelets which contain high numbers of MP’s and activated platelets have demonstrated that these are safe and effective in thrombocytopenic patients30 and in cardiac surgery patients.31

This study focused on characterising the in vitro haemostatic function of platelets in CS-WB. It is important to note that other aspects of platelet function, particularly their role in inflammation and immune responses, 32 have not been addressed here, partly as the consequences of these factors in transfused components is not fully understood. In this study we did not assess CS-WB using global assays (like thromboelastography, thrombin generation), or more physiological assays such as microfluidics. However, previously data from our group9 and others 10, 11, 12, 13, 33 have demonstrated that the ability to generate and maintain clot strength in CS-WB is maintained well up to 14 days with a slow reduction beyond that. This suggests that platelet funtion in CS-WB may be sufficient to generate clot formation, and the changes to activation and aggregatory response demonstrated in other assays may not have significant affects on the overall function. Furthermore, cold-stored platelets have been demonstrated to have increased aggregation under shear, which may have further benefit in major haemorrhage. 34, 35

For platelets there are no agreed standards that can be used to set specific limits, in part because there is little understanding how these measures relate to function in vivo and in part because of significant lab to lab variation in assays. The only measure for platelets where there is an agreed limit is pH where we know that for platelets stored in plasma (but not PAS), a pH of 6.2-6.4 correlates with reduced platelet function. 18, 36 However, this is not a useful specification for whole blood as red cells will also contribute to changes in pH. Most blood operators internationally apply a variety of assays to establish effect on platelet function in the laboratory prior to clinical studies. For these reasons it is very hard to consider parameters in isolation and set absolute limits. Schlicter et al demonstrated that platelets in CS-WB exhibit function for a minimum of 6 hours post-transfusion for up to 12 days of non-agitated storage, and up to 15 days with end-over-end agitation, suggesting this is enough to keep a critical patient stable until surgery. 22 The data from our study presented here, together with previous published work on CS-WB9, indicate that a shelf life of up to 21 days for CS-WB would be acceptable for use in clinical practice. However, large prospective randomised trials of CS-WB versus component therapy are needed to confirm the in vivo effect of the laboratory changes seen over time with CS-WB.

Acknowledgements

We thank the Component Development Lab for their help with the laboratory study especially Chloe Cavagnetto and Jenny Jolley who helped with assay validations, and the Hospital Services team in Cambridge for their assistance with selection of units. We also thank the Whole Blood programme team for all their contributions, and the London Air Ambulance and Barts Charity for the financial support of this project.

Reference List

- Beutler, E., West, C, (1979) The storage of hard‐packed red blood cells in citrate‐phosphate‐ dextrose (CPD) and CPD‐adenine (CPDA‐1). Blood, 54: 280‐ 4.

- Dawson, R.B., Ellis, T.J., Hershey, R. T., (1976) Blood preservation XVI packed red cell storage in CPD‐adenine. Transfusion, 16: 79‐ 81

- Holcomb, J.B., Tilley, B.C., Baraniuk, S., et al (2015) Transfusion of plasma, platelets, and red blood cells in a 1:1:1 vs a 1:1:2 ratio and mortality in patients with severe trauma: the PROPPR randomized clinical trial. JAMA, 313: 471-482.

- Jones, A. R. & Frazier, S. K. (2014) Increased Mortality in Adult Patients With Trauma Transfused With Blood Components Compared With Whole Blood. Journal of Trauma Nursing, 21, 22-29.

- Hunt, B. J., Allard, S., Keeling, D., et al (2015) A practical guideline for the haematological management of major haemorrhage. Br. J. Haematol. 170: 788-803.

- Nessen, S. C., Eastridge, B. J., Cronk, D., et al (2013) Fresh whole blood use by forward surgical teams in Afghanistan is associated with improved survival compared to component therapy without platelets. Transfusion, 53 Suppl 1, 107s-113s.

- [NG39], N.g. (2016) Major trauma: assessment and initial management. https://www. nice. org. uk/guidance/ng39/chapter/Recommendations-for-research

- Lozano, M.L., Pérez‐Ceballos, E., Rivera, J., et al (2003), Evaluation of a new whole‐blood filter that allows preparation of platelet concentrates by platelet‐rich plasma methods. Transfusion, 43: 1723-1728

- Huish, S., Green, L., Curnow, E., et al (2019), Effect of storage of plasma in the presence of red blood cells and platelets: re‐evaluating the shelf life of whole blood (Editorial, p. 3293). Transfusion, 59: 3468-3477

- Wang, S., Zhao, G., Li, N., et al (2019), An in vitro study of coagulation properties in refrigerated whole blood and reconstituted whole blood. Vox Sang, 114: 694-700.

- Meledeo, M.A., Peltier, G.C., McIntosh, C.S., Bynum, J.A. and Cap, A.P. (2019), Optimizing whole blood storage: hemostatic function of 35‐day stored product in CPD, CP2D, and CPDA‐1 anticoagulants. Transfusion, 59: 1549-1559.

- Sivertsen, J., Braathen, H., Lunde, T. H. F., et al (2018) Preparation of leukoreduced whole blood for transfusion in austere environments; effects of forced filtration, storage agitation, and high temperatures on hemostatic function. J Trauma Acute Care Surg. 84:S93-S103

- Remy, Kenneth, E., Yazer, Mark, H., Saini, Arun, et al (2018), Effects of platelet-sparing leukocyte reduction and agitation methods on in vitro measures of hemostatic function in cold-stored whole blood. Journal of Trauma and Acute Care Surgery, 84: S104–S114

- Reddoch, K. M., Pidcoke, H. F., Montgomery, R. K., et al. (2014) Hemostatic function of apheresis platelets stored at 4 °C and 22 °C. Shock; 41:54-61.

- Getz, T. M., Montgomery, R. K., Bynum, J. A., et al (2016), Storage of platelets at 4°C in platelet additive solutions prevents aggregate formation and preserves platelet functional responses. Transfusion, 56: 1320-1328.

- Nair, P.M., Pandya, S.G., Dallo, S.F., et al (2017), Platelets stored at 4°C contribute to superior clot properties compared to current standard‐of‐care through fibrin‐crosslinking. Br J Haematol, 178: 119-129.

- Murphy, S., Gardener, F. H. (1969) Effect of storage temperature on maintenance of platelet viability--deleterious effect of refrigerated storage. New England Journal Medicine. 280: 1094-1098.

- Murphy, S., & Gardner, F. H. (1971). Platelet storage at 22 degrees C; metabolic, morphologic, and functional studies. The Journal of clinical investigation, 50(2), 370–377.

- Valeri, C.R. (1976), Circulation and Hemostatic Effectiveness of Platelets Stored at 4 C or 22 C:. Transfusion, 16: 20-23

- Becker, G.A., Tuccelli, M., Kunicki, T., Chalos, M.K. and Aster, R.H. (1973), Studies of Platelet Concentrates Stored at 22 C and 4 C. Transfusion, 13: 61-68

- Filip, D.J., Aster, R.H., Relative hemostatic effectiveness of human platelets stored at 4 degrees and 22 degrees C (1978). J Lab Clin Med. Apr;91(4):618-24.

- Slichter, S. J., Fitzpatrick, L., Osborne, B., et al (2019) Platelets stored in whole blood at 4°C: in vivo posttransfusion platelet recoveries and survivals and in vitro hemostatic function. Transfusion. 59:2084-2092.

- Apelseth, T. A., Kristofferson, E. K., Kvalheim, V. L., et al (2017) Transfusion with Cold Stored Platelets in Patients Undergoing Complex Cardiothoracic Surgery with Cardiopulmonary Bypass Circulation: Effect on Bleeding and Thromboembolic Risk, Transfusion, Vol. 57, 3S: 3.

- Blajchman, M. A., Lee, D.H., (1997) The thrombocytopenic rabbit bleeding time model to evaluate the in vivo hemostatic efficacy of platelets and platelet substitutes, Transfusion Medicine Reviews, 11: 95-105.

- Strandenes, G., Sivertsen, J., Bjerkvig, C. K., Fosse, T. K., Cap, A. P., Del Junco, D. J., Kristoffersen, E. K., Haaverstad, R., Kvalheim, V., Braathen, H., Lunde, T. H. F., Hervig, T., Hufthammer, K.O., Spinella, P.C., Apelseth, T.O. (2020) A Pilot Trial of Platelets Stored Cold versus at Room Temperature for Complex Cardiothoracic Surgery. Anesthesiology. 133(6):1173-1183.

- Manno, C. S., Hedberg, K. W., Kim, H. C., Bunin, G. R., Nicolson, S., Jobes, D., Schwartz, E., Norwood, W.I., (1991) Comparison of the Hemostatic Effects of Fresh Whole Blood, Stored Whole Blood, and Components After Open Heart Surgery in Children. Blood, 77(5): 930-936.

- Yazer, M. H., Glackin, E. M., Triulzi, D. J., et al (2016), The effect of stationary versus rocked storage of whole blood on red blood cell damage and platelet function. Transfusion, 56: 596-604

- Aleman, M. M., Gardiner, C., Harrison, P. et al (2011), Differential contributions of monocyte‐ and platelet‐derived microparticles towards thrombin generation and fibrin formation and stability. Journal of Thrombosis and Haemostasis, 9: 2251-2261.

- Van Der Meijden, P. E., Van Schilfgaarde, M., Van Oerle, R., et al (2012), Platelet‐ and erythrocyte‐derived microparticles trigger thrombin generation via factor XIIa. Journal of Thrombosis and Haemostasis, 10: 1355-1362.

- Slichter, S. J., Dumont, L. J., Cancelas, J. A., et al (2018), Safety and efficacy of cryopreserved platelets in bleeding patients with thrombocytopenia. Transfusion, 58: 2129-2138

- Reade, M.C., Marks, D.C., Bellomo, R., et al (2019), A randomized, controlled pilot clinical trial of cryopreserved platelets for perioperative surgical bleeding: the CLIP‐I trial (Editorial, p. 2759). Transfusion, 59: 2794-2804

- Cognasse, F., Laradi, S., Berthelot, P., et al (2019). Platelet Inflammatory Response to Stress. Frontiers in immunology, 10, 1478. doi:10.3389/fimmu.2019.01478

- Pidcoke, H. F., McFaul, S. J., Ramasubramanian, A.K., et al (2013) Primary hemostatic capacity of whole blood: a comprehensive analysis of pathogen reduction and refrigeration effects over time. Transfusion, 53 Suppl 1:137S-149S.

- Montgomery, R.K., Reddoch, K.M., Evani, S.J., Cap, A.P. and Ramasubramanian, A.K. (2013), Enhanced SIPA in Refrigerated PLTs. Transfusion, 53: 1520-1530

- Six, K.R., Devloo, R., Compernolle, V. and Feys, H.B. (2019), Impact of cold storage on platelets treated with Intercept pathogen inactivation. Transfusion, 59: 2662-2671

- Solberg, C., Holme, S. and Little, C. (1986), Morphological Changes Associated with pH Changes during Storage of Platelet Concentrates in First‐Generation 3‐Day Container. Vox Sanguinis, 50: 71-77

Table 1 Final composition of blood components. Cold-stored whole blood (CS-WB) data includes CS-WB(NM) data as all were produced using Terumo Imuflex WB-SP. N= 20 Values = mean ± SD. All other components (Cold-stored platelet rich plasma (CS-PRP), standard room temperature platelet concentrates (RT-PLTS) and cold-stored platelet concentrates (CS-PLTS)) N= 10 Values = mean ± SD. *Data from assessment of Terumo WB-SP Imuflex filtered WB prior to separation of platelet rich plasma.

|  | CS-WB (inc. CS-WB(NM) | CS-PRP | RT-PLTS | CS-PLTS |
| --- | --- | --- | --- | --- |
| Volume (mL) | 470 ± 3 | 222 ± 38 | 301 ± 5 | 299 ± 10 |
| Platelet count per unit (x109/unit) | 77 ± 22 | 61 ± 21 | 271 ± 28 | 269 ± 29 |
| Platelet concentration (x109/L) | 158 ± 44 | 271 ± 76 | 900 ± 94 | 900 ± 94 |
| LD performance | ≥1 x 106/unit in 95 % of units | ≥1 x 106/unit in 90 % of units* | ≥1 x 106/unit in 100 % of units | ≥1 x 106/unit in 100 % of units |

Table 2: Comparison of platelet parameters in CS-WB to other platelet component

|  | Effect of storage in CS-WB | Comparison of CS-WB to CS-PLTS in this study | Comparison of CS-WB to RT-PLTS in this study | Comparison of CS-WB with other published work on CS-WB | Comparison of this study with other published work on CS-PLTS | Reference |
| --- | --- | --- | --- | --- | --- | --- |
| Platelet count | Decreased | Increased loss as % | Increased loss as % | Equivalent loss |  | Remy et al, 2018 Meledeo et al, 2019 |
| Lactate dehydrogenase | Small increase |  |  |  |  |  |
| Phosphatidylserine exposure | Highly increased | Greater | Much greater | Equivalent | CS-PLTS > RT-PLTS with lactadherin | Reddoch et al, 2014 |
| Platelet activation – P-selectin expression | Highly increased | Greater | Much greater | Equivalent | CS-WB > CS-PLTS > RT-PLTS | Reddoch et al, 2014, Sivertson et al, 2018, Slichter et al, 2019 |
| Response to ADP by Flow Cytometry | Increase over basal level decreased over time, maximal P-selectin expression maintained | Max response retained longer | Max. response retained much longer | No data available | No flow cytometry data available. Suggestion that max. response maintained longer during storage in CS-PLTS and CS-WB than RT-PLTS | No data available |
| Response to ADP by aggregometry | Decreased over time |  |  | Equivalent aggregatory response |  | Getz et al, 2016, Meledeo et al, 2019, Reddoch et al, 2014, Remy et al, 2018, Sivertson et al, 2018 |
| Response to TRAP by Flow Cytometry | Increase over basal level decreased over time, maximal P-selectin expression maintained | Max response retained longer | Max. response retained much longer | No data available | No flow cytometry data available. Suggestion that max. response maintained longer during storage in CS-PLTS and CS-WB than RT-PLTS | No data available |
| Response to TRAP by aggregometry | Decreased over time |  |  | Equivalent aggregatory response |  | Getz et al, 2016, Meledeo et al, 2019, Reddoch et al, 2014, Remy et al, 2018, Sivertson et al, 2018 |
| Response to CRP-XL by Flow Cytometry | Increase over basal level decreased over time, maximal P-selectin expression maintained | Max response retained longer | Max. response retained much longer | No data available | No flow cytometry data available Suggestion that max. response maintained longer during storage in CS-PLTS and CS-WB than RT-PLTS | Getz et al, 2016, Reddoch et al, 2014 |
| Microparticle | Increase |  |  | No data available | CS PLT/RT-PLT > CS-WB per unit | Reddoch et al, 2014, Wood et al, 2018 |
| Effect of non-agitation | No effect |  |  | Equivalent demonstration of no effect | No effect on in vitro parameters of cold stored platelets in CS-WB or concentrates | Sivertson et al, 2018, Yazer et al 2016. |

Figure legends:

Figure 1 Study design.

CS-WB was compared to standard NHSBT pooled buffy coat platelets in SSP+ held in current storage conditions (22±2 °C, agitation, RT-PLTS) or refrigeration (4±2 ° C, static, CS-PLTS). An additional comparator, platelets in plasma from a single LD whole blood donation (cold stored platelet rich plasma [CS-PRP]), were also used to understand the effect of storing CS-WB platelets at 4 °C in the absence of red cells. All units were stored and assessed across 21 days.

Figure 2 Platelet count, supernatant lactate dehydrogenase, annexin V binding and P-selectin expression of platelets during storage.

A) platelet count B) supernatant lactate dehydrogenase C) resting annexin V binding of platelets D) resting P-selectin expression on platelet surface in Cold-stored whole blood (CS-WB) (closed circles), cold-stored platelet rich plasma (CS-PRP) (closed triangles), standard room temperature platelet concentrates (RT-PLTS) (closed squares) and cold-stored platelet concentrates (CS-PLTS) (closed diamonds). N = 10. Values are mean ± SD. *CS-WB, CS-PRP, CS-PLTS and RT-PLTS were significantly higher than RT-PLTS at day 7 (current standard component at end of shelf life) (P<0.0001). #Annexin V binding in CS-WB is significantly greater than in CS-PLTS at days 14 and 21 (P<0.0001). There was no significant difference between CD62P expression in CS-WB and CS-PLTS at days 7, 14 and 21. Specific time-point comparisons were only made at days 7, 14 and 21 and were not made for A) and B).  ****Change over time in platelet count for CS-WB was significantly higher than RT-PLTS and CS-PLTS (P<0.0001), but not CS-PRP. Change over time in supernatant LDH levels over storage in CS-WB was not statistically different compared to RT-PLTS, CS-PLTS or CS-PRP.

Figure 3 Platelet activation and response to ADP, TRAP-6 and CRP-XL. P-selectin expression on platelets in cold-stored whole blood (CS-WB), standard room temperature platelet concentrates (RT-PLTS) and cold-stored platelet concentrates (CS-PLTS) was assessed without agonists (solid squares), and in response to maximal concentrations of ADP (2 µM), TRAP-6 (32 µM) and CRP-XL (2 µg/mL) – solid circles.  N = 10. Values are mean ± SD. *Change over time in maximal response was significantly lower in CS-WB than RT-PLTS for all three agonists (P<0.0001), and # significantly different between CS-WB and CS-PLTS with CRP-XL (P=0.0008).

Figure 4 Impedance aggregometry to ADP and TRAP-6. A) Cold-stored whole blood (CS-WB) response to ADP (HI 2 µM, Lo – 0.2 µM) in Impedance aggregometry. B) CS-WB response to TRAP-6 (HI 32 µM, Lo 4 µM). Cold-stored platelet concentrates (CS-PLTS) and standard room temperature platelet concentrates (RT-PLTS) response to C) ADP (2 µM) and D) TRAP-6 (4 µM). CS-WB was tested neat, whilst CS-PLTS and RT-PLTS were diluted to a fixed concentration of 600 x 109/L with autologous platelet-poor plasma each day. N = 10. Values are mean ± SD.

Figure 5 Platelet microparticle enumeration and functionality in CS-WB. CD61 positive, annexin V-binding microparticles in supernatants from cold-stored whole blood (CS-WB) were enumerated by flow cytometry. The ability for the supernatants to generate thrombin and produce a fibrin clot in phospholipid-depleted plasma was assessed in a commercially available kit. The dilution of plasma with SSP+ in standard room temperature platelets (RT-PLTS) and cold-stored platelet concentrates (CS-PLTS) meant it was not possible to measure activity in these units. N = 10. Values are mean ± SD.